## Slide 1
Development of a chromatography-free targeted DART- MS method for rapid detection of adulterants in saffron
Linda Monaci1*,
A. Luparelli1, L. Quintieri1, A. Verdu2
Institute of Biomembranes, Bioenergetics and Molecular Biotechnologies, (IBIOM-CNR), Bari, Italy.
Bruker France S.A.S. Applied Mass Spectrometry, Wissembourg Cedex, France;

## Slide 2
Coupling DART-MS:
A Powerful Tool for Rapid Screening
Direct Analysis in Real Time Mass Spectrometry (DART-MS) is a relatively recent technique (introduced in 2005) designed for rapid mass spectral analysis under ambient conditions.
✨Why is DART-MS so attractive?
Ionization | Ambient pressure; direct desorption/ionization (no vacuum chamber needed)
Speed | Fast MS fingerprint in seconds (Gross, 2014)
Simplicity | Only two core components: Ion Source + MS
No Chromatography | Chromatography-free → no separation, no long runs
Sample Prep | Minimal or none → ideal for rapid screening
Throughput | Compatible with multiwell sample cards for high-throughput workflows
- 2° Green Analytical Chemistry Workshop -
Notes: DARTmass spectrometry is the most significant discover in mass spectrometry since the invention of the electrospray ionization (ESI) and matrix-assisted laser desorption/ion ization (MALDI) techniques.

## Slide 3
How does it work?
Direct Analysis in Real Time Mass Spectrometry (DART-MS) is technique introduced in 2005, designed for rapid mass spectral analysis under ambient conditions.
Inventors:
Cody and Laramee
FAST
EASY SOLVENT FREE
IONIZATION: due to penning ionization and proton transfer
- 2° Green Analytical Chemistry Workshop -
Notes: A gas stream passes through the discharge region where ions, electrons and metastable (excited) molecules or atoms are produced. Cody and Laramee (inven tors of DART-MS) used the term `electrical discharge'source exit and the mass spectrometer inlet. Most DART MS experiments are carried out with helium (purity 99.999%).
It allows one to carry out a fast non-contact analysis of various samples. Most often, solid samples are introduced into the ionization region with a pincer while liquid ones are introduced with a glass or polymeric rod. Evaporation/desorption of analytes from the surface of samples in a helium or nitrogen stream is accompanied by their ionization involving metastable atoms or molecules (or reagent ions). Only the low-molecular-mass components are ionized; this allows one to avoid the impact of the high-molecular-mass matrix components of the analysis. It is of interest not only to identify analytes, but also to quantify them.
Technical part:
the potential at the gas-discharge needle of the DART source ranges from 1 to 5 kV, the first perforated disk electrode is grounded and the potentials at the second perforated electrode and at the grid electrode are about 100 and 250 V, The gas flow rate is usually at least 0.55 litre min 1 and the gas temperature varies from ambient temperature to 350 C.Thewidthof thegap betweenthe DARTsourceexit and the skimmer input orifice or a vacuum interface of a mass spectrometer usually varies from 5 to 25 mm
It is assumed that a number of complex processes occur concurrently depending on the proton affinity and the ionization potential of analyte, on the analyte concentration in the sample, on the polarity and nature of the gas used and on the presence of auxiliary reactants.

## Slide 4
The helium plasma, generated from an electrical discharge in a ceramic flow cell, is used in an ion source by applying a potential between a needle-electrode and a ground counter-electrode (in red).
The solvent-free stream of helium consists of excited neutral helium atoms, which ionize and generate protonated water molecules which vaporize the desired analytes by protonating the sample.
The ionization occurs when metastable atom transfers its energy to an analyte molecule. This takes place in the sample gap (circled in blue) and is due to penning ionization and proton transfer (metastable atom transfers its energy to an analyte molecule resulting in the formation of a molecular ion).
Quantitative capabilities of DART-MS to be further investigated
The ionization amount of a target analyte correlates with the molecular ion peak intensity.
How does it work?
- 2° Green Analytical Chemistry Workshop -

## Slide 5
Trend of Publications in the last decade
Keywords: «DART» AND «Mass Spectrometry»
2005-2025
50
- 2° Green Analytical Chemistry Workshop -

## Slide 6
Forensics
Environment
Food analysis
Production
Processing
Storage
Transportation
Food components
Food contaminants
Authenticity and traceability
Applications of DART
- 2° Green Analytical Chemistry Workshop -

## Slide 7
- 2° Green Analytical Chemistry Workshop -

## Slide 8
Research Background & joint collaboration
Tipically combined with HR MS platforms: Q TOF, TOF, Orbitrap MS for food analysis.
Rare reports about quantitative analysis in food by DART-MS technology.
The launch on the market of the DART TQ MS (EVOQ) has prompted in investigating instrumental capabability of TQ MS for quantitative analysis of contaminants/adulterants in foods.
- 2° Green Analytical Chemistry Workshop -

## Slide 9
Case Study
Food authenticity:
detection of adulterated saffron by low-cost spices
1
- 2° Green Analytical Chemistry Workshop -

## Slide 10
Commercial saffron is obtained from the dried stigmas of Crocus sativus L. (Iridaceae family) and it is currently classified as the most expensive spice in the world.
The costs for saffron production and harvesting are high, being consequently a common target for adulterations to increase profit and its history of adulteration dates back to about 600 years ago (A. Koocheki, E. Milani, 2020).
In Europe, saffron is sold:
as entire stigmas, considered the best quality product;
as powder
as an ingredient of commercial processed foodstuffs
In the latter forms (as powder and ingredient), the adulterations with other colorants or plant tissues can be easily performed.
SAFFRON
- 2° Green Analytical Chemistry Workshop -
Notes: Saffron (Crocus sativus) is a high value spice obtained from drying the stigmas of Crocus sativus L., mainly produced in India, Iran and other Mediterranean countries. It is classified as the most expensive spice with prices of $1,100 -11,000 per kg and therefore a prime target for fraudsters. The most common substances illegally added to saffron are powders from turmeric, safflower and paprika (or other Crocus species) as well as natural and artificial dyes including erythrosine, ponceau 4R, tartrazine, Sudan dyes, magenta III and rhodamine B, Safranal, which is a synthetic aroma compound.

## Slide 11
2
Current State of Fraud in Europe
(Maquet, A. 2021)
Food authenticity:
detection of adulterated saffron by low-cost spices
1
Timing: The coordinated control plan was executed in 2019-2021.
Samples: 1900 samples from Twenty-one EU Member States plus Norway and Switzerland were tested.
Methods used: High performance liquid chromatography coupled to high resolution mass spectrometry (HPLC-HRMS), energy dispersive X-ray fluorescence spectroscopy (ED-XRF), thermogravimetric analysis (TGA), digital droplet polymerase chain reaction (ddPCR), next generation sequencing (NGS), real-time PCR (rt-PCR); fourier transform infrared spectroscopy (FTIR)
RESULTS: The overall rate of suspicious samples was 17%
Case Study
- 2° Green Analytical Chemistry Workshop -
Notes: Culinary herbs and spices are valued for flavouring food and could also provide other beneficial properties, such as antioxidative and bacteriostatic effects as well as certain pharmacological activities. Their supply chain is complex, long, and globalised.
Europe is one of the world’s leading importing regions for herbs and spices, importing approximately 300.000 tons, mostly spices from East Asia. Most of the spices are produced in countries where certain post-harvest processes such as drying and cleaning may happen before being shipped to the importing country where they are further cleaned and sanitised before being packaged and distributed either to other food businesses or for retail consumption. At each stage, fraudulent manipulations may happen and the more often the material is transferred from one operator to the next, the fraud opportunity increases.
Information available to the European Commission indicates that adulterated herbs and spices are present on the EU market but remain undetected. Therefore, the European Commission set up a coordinated control plan in the period 2019-2021 inviting the EU member states to sample certain herbs and spices and send them for analysis to the Joint Research Centre. The main objective of the plan was to establish the prevalence on the market of some non-compliances and of some possible illegal practices into marketing of herbs and spices. Twenty-one EU member states plus Norway and Switzerland submitted nearly 1900 samples to JRC for analysis. The majority of samples was ground or crushed.
The percentage of samples which were suspicious of adulteration were 17% for pepper and 11% for saffron
The methodologies used include HRMS analysis, X ray florescence spectroscopy, ddPCR, RT-PCR and FTIR.
Different methods have been developed for detecting saffron adulteration. However, they all have difficulties to detect adulterants like safflower added to a level lower than 20% (w/w). Here, we developed a rapid, robust, high-throughput DART method for the reliable quantification of saffron adulteration with low safflower contents on a triple quad MS.

## Slide 12
Case Study
Food authenticity:
detection of adulterated saffron by low-cost spices
1
The aim of this study was to develop a fast, robust, and chromatography-free DART–MS/MS method for the targeted detection of turmeric and safflower adulteration in saffron.
- 2° Green Analytical Chemistry Workshop -

## Slide 13
Methods for detecting saffron adulteration although they are time consuming and complex. Such as gas chromatography (GC), thin layer chromatography (TLC), liquid chromatography-mass spectroscopy (LC-MS) , nuclear magnetic resonance (NMR), molecular methods (PCR).
The official ISO 3632-2 protocol, is the current standard method for saffron quality analysis, combines UV–vis spectroscopy with HPLC to determine color strength and crocin content. But the official method was unable to detect saffron adulterants (safflower, marigold, or turmeric) added to a level lower than 20% (w/w).
Food authenticity:
detection of adulterated saffron by low-cost spices
1
Case Study
Background and aim of the study
ISO 3632 remains the gold standard for
saffron quality
NEED TO DEVELOP A FAST AND GREEN METHOD FOR RAPID DETECTION OF ADULTERATED SAFFRON SAMPLES
- 2° Green Analytical Chemistry Workshop -
Notes: While ISO 3632 remains the gold standard for saffron quality, there is some debate about the reliability of the spectrophotometric method used to measure saffron properties. This method combines UV–vis spectroscopy and high-performance liquid chromatography (HPLC) to determine color strength and crocin content. This official method was unable to detect saffron adulterants (safflower, marigold, or turmeric) added to a level lower than 20% (w/w).
Other methods include electronic nose, GC.MS and microscopic analysis

## Slide 14
DART-MS/MS analysis
Add 5mL of EtOH/H₂O 70/30 (v/v) to 50 mg of spice powder
    Shaking 15 min
    Centrifuge 5min at 13000rpm
    Filter on PTFE 0,45uM
Serial dilution of saffron extract intentionally adulterated with increasing concentrations of safflower and turmeric extracts.
Extract deposition
Safflower 100%
Saffron 100%
Turmeric 100%
Quick extraction
Materials and Methods
Food authenticity:
detection of adulterated saffron by low-cost spices
Case Study
- 2° Green Analytical Chemistry Workshop -

## Slide 15
Materials and Methods
Identification of potential markers tracing for safflower and turmeric in saffron
Food authenticity:
detection of adulterated saffron by low-cost spices
Calibration curves in saffron adulterated with safflower or turmeric powder
Application of the developed multi-target method to adulterated saffron samples
STEP 1
STEP 2
Identification of ion transitions for each adulterating compound
STEP 5
STEP 3
Optimization of DART and MS settings for MRM-DART TQ MS/MS analysis
STEP 4
safflower
turmeric
DART-TQ parameters
- 2° Green Analytical Chemistry Workshop -

## Slide 16
Scanning mode: Linear/ Circular scanning/ jumpshot
Other parameters: scanning speed, scanning range
Ion source parameters:
Grid/cone Voltage (40V)/
Cone temperature (350°C)
cone gas pressure (20)
DART temperature (300°C)
Gas type helium
Gas time program: 6 seconds
DART MS parameters
Results
Food authenticity:
detection of adulterated saffron by low-cost spices
- 2° Green Analytical Chemistry Workshop -

## Slide 17
Full MS ion scan
Once the precursor ions had been identified….
Product Ion scan experiments on the selected precursor ions
Optimization of CE to improve fragments detection (5,10,25,50 eV)
Identification of best transitions and MRM method set up
Introduction of IS and repeatability check

## Slide 18
Food authenticity:
detection of adulterated saffron by low-cost spices
Selection of Safflower and Turmeric Markers
Full scan MS spectra obtained in the mass range 100–600 m/z in pure saffron (A), safflower (B) and turmeric (C).
Results
- 2° Green Analytical Chemistry Workshop -

## Slide 19
Final conditions for safflower/turmeric detection in saffron
Results
Experiments based on product ion-monitoring mode were directed to optimize the collision energy among the values tested 5, 10, 15, 20, and 40 eV, confirming that 10 eV was the best compromise in terms of fragment abundance and signal stability over repetitive analysis as precursor/fragment ion ratio.
Candidate markers for turmeric to be further investigated:
Candidate markers for safflower to be further investigated:
    116,2
    198,2
    296,2
    446,4
    235,3
    309,0
    338,9
    369,0
Food authenticity:
detection of adulterated saffron by low-cost spices

## Slide 20
Full MS ion scan
Once the precursor ions had been identified….
Product Ion scan experiments on the selected precursor ions
Optimization of CE to improve fragments detection
Identification of best transitions and MRM method set up
Introduction of IS and repeatability check

## Slide 21
+10% safflower
+25% safflower
+50% safflower
Product ion scan [ion 116.2] at CE 10
Precursor ion overlay
selected marker 116.2
Saffron spiked with safflower at 50%
Saffron sample
Saffower extract
Results
Food authenticity:
detection of adulterated saffron by low-cost spices
Product ion scan of the selected precursor markers allowed to identify transitions 116/70 and 446/116 to be used for the MRM method
- 2° Green Analytical Chemistry Workshop -

## Slide 22
IS:
To compensate the errors and variability of the measurement (due to ionization processes /sample position or MS related factors)
Express the value as analyte-to-internal standard ion ratios rather than absolute intensities of the analyte signals.
Issues when the analyte and the internal standard have different boiling points. Since their ionization will occur asynchronously and this may introduce errors.
The most suitable IS should be sought for each method of analysis
Results
Fortified saffron extracts and calibration curves -
Introduction of an internal standard
Calculation of method repeatability

## Slide 23
Candidate markers for turmeric to be further investigated:
Candidate markers for safflower to be further investigated:
    116,2
    198,2
    296,2
    446,4
Food authenticity:
detection of adulterated saffron by low-cost spices
Results
50% and 20%
safflower fortification, 10 replicates of analysis
CV% respectively 11% and 17% without correcting with IS and 6 and 8% by peak area correction with IS.
Repeatability
Intermediate precision
five independent replicated analyses of saffron adulterated at 50% during three consecutive days with a calculated CV of 15% with IS correction.
Fortified saffron extracts and calibration curves
5 fortification levels 4 replicates each
Linearity
Quantifier ion transition (safflower): 116.2/70.1 (m/z)
with IS
no IS
R2 = 0, 92
R2 = 0, 98
- 2° Green Analytical Chemistry Workshop -
Notes: DART- MS/MS analysis were carried out on blank and saffron extract artificially adulterated with safflower extract. 5 fortification levels: 10 - 20 - 30 -50 - 70 % (of safflower in saffron); 4 technical replicates each

## Slide 24
Candidate markers for turmeric to be further investigated:
Candidate markers for safflower to be further investigated:
Food authenticity:
detection of adulterated saffron by low-cost spices
Results
Production of Matrix-Matched Calibration Curves in Saffron Powder
Fortification | Calibration curve equation | R² | LOD % (w/w) | LOQ % (w/w)
Safflower in saffron | Safflower in saffron | Safflower in saffron | Safflower in saffron | Safflower in saffron
116/70 | y = 0,0717 (±0,0063) x - 0,1786 (±0,174) | 0,985 | 7 | 24
446/116 | y = 0,1014 (±0,0067) x - 0,2633 (± 0,185) | 0,991 | 5 | 18
Turmeric in saffron | Turmeric in saffron | Turmeric in saffron | Turmeric in saffron | Turmeric in saffron
339/178 | y = 1,4027 (± 0,052) x - 3,5298 (± 1,445) | 0,997 | 3 | 10
368/285 | y = 4,2545 (± 0,253) x- 5,9038 (± 0,965) | 0,993 | 5 | 16
- 2° Green Analytical Chemistry Workshop -

## Slide 25
Candidate markers for turmeric to be further investigated:
Candidate markers for safflower to be further investigated:
    116,2
    198,2
    296,2
    446,4
Food authenticity:
detection of adulterated saffron by low-cost spices
Results
Recovery Assessment,
Recovery experiments were performed on three replicated samples at 20% level each, and recoveries were in the range 53-66%
53 ± 4% monitored marker 339 (tracing for turmeric)
66 ± 5% monitored marker 446 (tracing for safflower)
- 2° Green Analytical Chemistry Workshop -
Notes: Analytical GREEnness calculator, a comprehensive, flexible, and straightforward assessment approach that provides an easily interpretable and informative result.

## Slide 26
Greeness assessment
Food authenticity:
Results
The assessment criteria are taken from the 12 principles of green analytical chemistry (SIGNIFICANCE) and are transformed into a unified 0–1 scale. The final score is calculated based on the SIGNIFICANCE principles.
https://mostwiedzy.pl/AGREE.
- 2° Green Analytical Chemistry Workshop -
Notes: Inorder to comply with recommendations of the green analytical chemisty and in order to promote a method environmentally sustainable we calculated the greenness of the method by applying a software recently developed by scientists. This a analytical GREEnness calculator, a comprehensive, flexible, and straightforward assessment approach that provides an easily interpretable and informative result.
With the aim of comprehensively evaluating the greenness of analytical methodologies, we have converted each of the 12 principles of GAC into scores. Criteria taken into consideration are generation of waste, reagents that are persistent, bioaccumulative, or toxic, whether reagents are hazardous, whether the conditions are corrosive.

## Slide 27
Results
A fast, robust, and chromatography-free DART–MS/MS method has been developed for the targeted detection of turmeric and safflower adulteration in saffron.
- 2° Green Analytical Chemistry Workshop -

## Slide 28
Application note recently published

## Slide 29
Lorem ipsum dolor sit amet, consectetur adipiscing elit, sed do eiusmod tempor incididunt ut labore et dolore magna aliqua. Ut enim ad minim veniam, quis nostrud exercitation ullamco laboris nisi ut aliquip ex ea commodo consequat.
Strengths and Limitations of the DART-MS method
 | Advantages | Disadvantages / Limitations
Sensitivity | ✅ High sensitivity for both adulterants and contaminants | ❌ Potential matrix effects in complex samples
Speed | ✅ Real-time analysis (no chromatography required) | ❌ Limited separation of isomeric compounds
Sample prep | ✅ Minimal to no sample preparation required | ❌ Not always suitable for very fatty or highly complex matrices
Versatility | ✅ Applicable to both adulteration markers and chemical contaminants | ❌ Requires individual optimization for each new analyte or matrix
Specificity | ✅ MRM allows high selectivity for targeted compounds | ❌ Restricted to compounds with known diagnostic transitions
Scalability | ✅ Potential for high-throughput screening | ❌ Extensive validation needed for regulatory implementation
- 2° Green Analytical Chemistry Workshop -

## Slide 30
Conclusions
The method is rapid, high-throughput, and environmentally friendly, requiring no chromatographic separation and allowing analysis in just 30 seconds per sample.
The DART-TQ MS/MS multi-target method successfully detected saffron adulteration down to 3% turmeric and 5% safflower.
The inclusion of an internal standard significantly contributed to decrease variability of the MS measurements.
extend the method to other potential adulterants to deploy a single multi-target method for the detection of multiple substances
FUTURE PERSPECTIVES:

## Slide 31
Thank you
CNR
Anna Luparelli
Laura Quintieri
WIlliam Schirinzi
BRUKER
Alexandre Verdu
Noud van den Borg
Elena Bueno
Giuseppe Labella